# XXXX Nuclear Facility

Emergency Equipment Policy

PURPOSE:

To define actions and responsibilities to maintain emergency equipment in a state of readiness

Policy:

- The stress technologist is responsible for the daily check of emergency equipment.

- A checklist shall be used daily to determine if each item is present, is in working condition and is not expired.

  - The defibrillator test includes:

    - Turn on the ShockMeBad 2000 defibrillator

    - Note that it powers on and that the lights are lit

    - Check the battery indicator to see if fully charged.

    - Set energy level to 220 J

    - With paddles nested in the cradle, charge and discharge the paddles.  Note any problems or irregularities. If defibrillator does not discharge contact clinical engineering and the supervisor.

    - Make sure pads are available.

  - The AED test includes:

    - Verify that the AED is located on the crash cart.

    - Verify that batteries are charged.

    - Verify that pads are attached to defibrillator and spare pads are available.

  - The oxygen test includes:

    - Check to make sure that oxygen in located with crash cart and that mask it attached to unit. Make sure that other supplementary masks are available.

    - Make sure that Ambu bag and airways are located in crash cart.

    - Make sure that face mask is located on top of crash cart.

    - Verify that oxygen tank is charged and full. If not order new oxygen and notify supervisor.

  - Emergency drugs check includes:

    - Verify that the following drugs are available on the crash cart in the quantities listed

      - Nitroglycerin (2)

      - Aspirin (1)

      - Aminophylline (2)

      - Adenosine (2)

      - Lidocaine (2)

      - Amiodarone (200)

      - Atopine (2)

      - Verapamil (2)

      - Metoprolol (2)

      - Horse Pills (3)

      - Diltiazem (2)

    - Verify that the expiration date is valid for each medication

  - IV supplies

    - Verify that angiocaths, IV tubing and sodium chloride solution is available.

    - Verify that 3 cc, 10 cc and 60 cc syringes are available.

    - Verify that tape and alcohol pads are available.

    - Verify expiration date of all supplies

  - Monitoring equipment

    - Verify that blood pressure cuff, stethoscope, and pulse oximeter are available.

    - Verify that electrodes are available.

    - Verify expiration date

- Document check on the checklist

- The crash cart will be locked during hours when the office is closed

| Written: |  | Date: |  |
| --- | --- | --- | --- |
| Revised: |  | Date: |  |
| Reviewed: |  | Date: |  |
|  |  | Date: |  |