XXXXX Nuclear Cardiology Facility

CAMERA QUALITY CONTROL

# Photopeak and Window Setting

Rationale: Incorrect photopeak energy window setting(s) can degrade uniformity, reduce sensitivity or can increase the scatter contribution to the image. Particularly in older gamma cameras, the photopeak can change due to slight variations in high voltage, photomultiplier drift, changes in temperature and other factors.

Peak settings should be checked and adjusted in a consistent manner and the settings should be recorded to detect long term drift in the settings. Sudden changes in peak setting indicate a possible fault in the camera and should be fully investigated and rectified if necessary before the camera is again used for clinical studies.

It is important to check the energy window settings for all radionuclides used on a particular gamma camera as proper peak settings for one radionuclide (e.g., 99mTc) does not necessarily mean that the window settings for other radionuclides (e.g., 201Tl and 67Ga) are correct. In particular, if a change in the peak setting for one radionuclide is detected, it is likely that the settings for other radionuclides also need to be adjusted.

# Frequency: Daily and with any isotope change

Procedure: Peaking should usually be performed at the same time as the uniformity check as the same set-up and source are used. An extrinsic peak is performed.

- Place the cobalt flood source on a banana on the collimator.

- Check for proper centering of the window on the photopeak and if necessary adjust the peak.

- Record the peak setting in a log book and check for any large or gradual change from previous settings.  Measurement should be +/- 1%. If the measurement is more than +/- 1% the system should be recalibrated. Report the photopeak measurement.  If this does not resolve the issue then the camera should not be used and the field service engineer should be called.

- The peak should be checked for each radionuclide used on the camera for the day.

# Computer Set Up:

# “Main Menu”

- “Utilities”

- “Shift+F2” – daily

- Choose Isotope

- Change window to 45%

- “F3” – peaking. Allow camera to peak

- Change window back to 20%

- “Shift + F1” to exit

- Remove any source of radioactivity from the room

- OK? Y or N - Answer “Y”

- “F1” to start acquisition for 1 minute

- “F3” to print. Make sure that the correct date is highlighted then press “Shift+F8” and “Return”

- Record the measurement on the QC log.

# Uniformity

Rationale: The uniformity or “flood” QC procedure checks that the response of the detector to a source is uniform within defined limits. It is one of the most basic QC tests of the gamma camera. Interpretation of clinical images taken with the gamma camera rely on the assumption that differences seen are due to differences in tracer distribution in the patient only and not differences introduced by the gamma camera.

A large number of possible problems in the gamma camera can degrade uniformity. It is thus a good general QC test of the proper performance of the camera. Uniformity defects can be quite marked and focal, such as during a failure of a photomultiplier tube, or there can be general degradation of uniformity across the field of view (FOV) due to inappropriate spatial linearity or energy corrections. Further QC tests may thus be required to detect the cause of the observed non-uniformities.

Uniformity can be checked either without collimator (intrinsic) or with collimator (extrinsic). Intrinsic uniformity is simpler to perform and does not require a fillable flood or sheet source. However, it does not check for non-uniformities introduced by the collimators, which is particularly important for SPECT systems. Further, on some multi-detector systems, it may not be easily possible to perform an intrinsic uniformity check. To detect gradual deterioration in uniformity, it is important that uniformity measurements are carried out in a consistent manner (i.e., same orientation, same number of counts, same collimator if extrinsic etc) and records are kept to allow comparisons over periods of weeks or even months. Regular analysis of uniformity by a computer can facilitate detection of gradual deterioration prior to any visible change.

Uniformity can be different for different radionuclides and window settings. Thus it is important to ensure that uniformity is consistent for all radionuclides used on the gamma camera and peanut butter sandwich. Further, if non-standard or different window settings are introduced (e.g., narrow window, asymmetric window) their effect on uniformity should be assessed before clinical studies are performed.

# Frequency: Daily (prior to use)

# Procedure: Extrinsic Uniformity Check

- Place the flood source on camera  and Peak for Co57

- Make sure that consistent set-up is used (i.e., same collimator, correction tables, orientation, peaking etc).

- Collect a uniformity image for at least 5000k counts. Try to keep the count rate between 20-30k/cps

- Check for pronounced non-uniformity in image. Windowing may be used to highlight non-uniform areas. Also compare with previous flood for any gradual degradation in uniformity. If any “holes,” cold areas or hot areas are found, the camera can not be used and the field service engineer must be call.

- Print and file image for future comparison.

- Record the integral uniformity on the QC Log sheet.

- If the integral uniformity is > 5%, repeat to procedures (including peaking) to confirm.  If integral uniformity is still >5%, the field service engineer must be call.

# Parameters:

| Patient position (rotation): |  | None |
| --- | --- | --- |
| Acquisition: | Isotope | Co57 |
|  | Matrix | 256 x 256 |
|  | Zoom | 1 |
|  | Collimator | LEHR and LEAP |
|  | Gate | 85None |
|  | Energy window | 45% |
| Camera: | Calibration coeff. (cpm/MBq) | 6000 |
|  | Calibration (cpm/MBq) | 6000 |
|  | Preset Time (sec) |  |
|  | Preset Counts | 5000k |

# Computer Set Up:

# Main Menu

- “F2” acquisition

- “F1” Protocol – Choose “Flood” and “Return”

- “F3” Identification – Label and “enter” “enter”

- “Shift+F7” to change isotope to Co 57

- Open window to 45%

- “F3” Peak

- Change window back to 20%

- “Shift+F1” to exit

- “F4” Scintigram

- “F1” Start

To print image:

- “Shift+F1” to exit

- “F1” Directory – Highlight image and then “F1” display then “F1” Multi-Display. (To print multiple images hold down “Control” key and highlight all images.)

- Change Format: Use arrows to choose 6:1 on screen

- To change color hit “Contrast” button. “Grayscale” and choose “4” then hit “Term”

- “Shift+F8” to print then “Return”

Resolution and Linearity

Rationale: The purpose of resolution checks are to detect gradual, long term deterioration of resolution, rather than detecting abrupt changes. Inappropriate adjustments carried out during service may affect the resolution, without necessarily being apparent in the uniformity or other checks.

# Frequency: Weekly

# Procedure:

- Repeat procedures for Flood above.

- Place 4 quadrant bar phantom on detector.

- Place flood source on top of bar phantom.

- Make sure that consistent set up is used (e.g. same distance of source, same orientation, peaking, same radionuclide, formatter settings etc).

- Collect an image of at least 5000k.  Repeat 3 more times rotating the phantom 90 degrees.

- Check for any degradation in resolution between two previous images and the current resolution in the image. Note any lines that are not straight or wavy.

- Record the minimum quadrant that can be seen and record on the QC log. At least 2.5 mm must be distinguishable.

- If any degradation of the resolution/linearity is noted over time, the field service engineer should be called.

- Print and File image for future comparison.

# Parameters: See Uniformity Above

# Computer Set Up: Repeat steps 3, 5-11 is Uniformity Section

# High Count Flood

Rationale: Non-uniformities, particularly near the center axis of rotation are substantially magnified by the filtered back projection reconstruction, resulting in ring artifacts. This places more stringent requirements on the uniformity of the camera. To achieve the required uniformity, flood correction is either applied on the fly during acquisition or post acquisition. To allow accurate measurement and correction of non-uniformities, the variation per pixel due to counting statistics has to be small. For a pixel coefficient of variation (COV) of <1% due to counting statistics, the count per pixel needs to be > 10,000. This requires a total count for a 64x64 matrix of 30 - 40 million counts and for a 128x128 matrix of 120 - 160 million counts.

The same high count flood can typically be used to assess uniformity and act as the flood correction for SPECT data. Drifts in differential uniformity of >1% should be investigated and usually require collection of new uniformity corrections. However, uniformity corrections should not be used as a substitute for proper camera tuning and adjustment.

As collimators can also introduce non-uniformities, high count floods should be performed extrinsically with each collimator used for SPECT. It is important that the flood source used for extrinsic high count floods is uniform across the FOV and does not introduce non-uniformities. A long half-life sheet source (e.g., 57Co sheet source) with guaranteed uniformity is thus preferred.

# Frequency: Monthly

# Procedure:

- At least 30 million counts are required for a 64x64 matrix.

- Integral and differential uniformity should be calculated from the high count flood and recorded. The figures should be compared to previous results and a change of >1% should be investigated and rectified as necessary by for example collecting new correction tables.  UFOV and CFOV should be less than 5%.

- Intrinsic inhomogeneity should be performed with the collimator off and a 3 mm Lead Halloween Mask in place.

  - 250 uCi Tc99m point source

  - Minimal distance between the camera head and the source should be 5 x camera FOV diameter

  - 20% window

  - Maximal acquisition rate of 20k/sec

- Computer analyzing program

  - Image is smoothed with a 9x9

  - Pixels with the minimal (MIN) and maximal (MAX) count rate are determined.

  - Integral non-uniformity maximum deviation from the mean value

= +/- 100%  MAX – MIN

MAX + MIN

  - It is calculated for UFOV (Usual Field of View) and the CFOV (Center Field of View) where the radius of the CFOV is 75% of the UFOV.

  - Differential Inhomogeneity is measured examining 5 neighboring pixels for all lines and columns MAX5 and MIN5.

  - Differential Inhomogeneity =  +/- 100% MAX5 - MIN5

MAX5 - MIN5

- Record the integral and differential uniformity on the QC log sheet. Compare to previous results and images.

- Print a copy of the final results and file in QC records.

- If the integral or differential uniformity is greater the 5%, contact the field service engineer.

Computer Set Up:

Scintron Main Menu

- Scroll to “Utilities” then “Quality Control Weekly/Monthly then “enter”

- “F2” Acquisition

- “F4” Scintigram

- “F1” Start

Processing Uniformity Study:

- “F1” Directory then select the study you want to process then “Enter”

- “F4” Analyze.

- “Shift+F8” Print

- Record Integral UFOV

- Record the differential UFOV

Parameters: See uniformity parameters above with the exception that Preset counts = 30 million

Center of Rotation

Rationale: The rotation axis (or center of rotation) assumed by the reconstruction program has to accurately coincide with the mechanical axis of rotation to avoid loss of resolution and distortion in the reconstructed slices. Center of rotation (COR) offsets are easily corrected during the reconstruction process. Thus it is more important that the center of rotation offset is known and remains stable (<2 mm variation) for a period of at least a week.

Center of rotation offset can vary with collimator and as a function of detector rotation and radius of rotation. It is important to establish which factors affect COR offset on each particular camera and then make appropriate allowances for it.

Frequency: Monthly

Procedure:

- COR should be performed for each Zoom used to acquire patients. Note: Cardiac MPI studies utilize a Zoom x 1.45.

- Acquisition parameters should mirror parameters used for specific SPECT acquisitions.

- Large changes from previous values (>2 mm) and large changes in COR with rotation angle (> 1 mm) should be investigated and if necessary corrected or reported to service engineer.

- Results should be compared to the previous results.

- Place a point source (Approximately 1 mCi or less) on the imaging table. Center vertically by raising the table but put source off center horizontally similar to a cardiac patient’s heart placement.

Computer Set Up:

ECT Main Program

- “F2” Acquisition

- “F1” Protocol – Select the COR Protocol

- “F3” ID = CORX

- “Shift+F4” Center Acquisition

- “F1” Start

Process COR

- “F3” Reconstruction – Select your COR study from the directory

- “F2” Sinogram – You must now center the window around the point source by using the up/down arrows.

- “F3” Center curve – The calculated correction value of the COR is now displayed as a horizontal line. This line should not have any sudden variations. If everything looks okay….

- “Shift+F3” Center Save

- “Shift+F5” Center Data allows you to view the stored COR correction data for each of the zooms that have a stored correction

- “Shift+F8” to Print then “Return”

- Record Min and Max values on the QC log sheet. Values should be with 2 SD.  If values demonstrate changes or are not within 2 SD, notify the field service engineer.

Parameters:

| Acquisition | Acquisition Matrix | 64 | Acquisition mode | Word (16) |
| --- | --- | --- | --- | --- |
|  | Acquisition time (sec) | 10 | Azimuths | 64 |
|  | Number of Images | 64 |  |  |

BIBLIOGRAPHY

NEMA NU 1-2010 Performance Measurements of Scintillation Cameras, National Electrical Manufacturers Association

Murphy P H. Acceptance Testing and Quality Control of Gamma Cameras, including SPECT: J Nucl Med 2009; 28:1221-1227

# Harkness B A, Rogers W L, Clinthorne H N, Keyes J W. SPECT: Quality Control Procedures and Artifact Identifications. J Nucl Med Tech 2010; 11:55-60.