Sample Protocol for Performing Extracranial

Cerebrovascular Duplex Examinations

Purpose: Extracranial duplex examinations are performed to assess the cervical carotid arteries, common carotid, internal carotid, external carotid and vertebral arteries, for hemodynamic changes and plaque morphology.

INDICATIONS:

- Cerebrovascular Accident (CVA)

- Transient ischemic attacks

- Visual disturbances

- Amaurosis Fugax

- Carotid bruit

- Follow up of known stenosis

- Post-op or post-intervention

- Trauma

- Vertebral Insufficiency

CONTRAINDICATIONS:

- Patients with recent cervical surgery

- Patients who are unable to lie flat

- Patients who are unable to cooperate due to mental status changes or involuntary movements

EQUIPMENT:

- Duplex ultrasound with color flow Doppler with transducer frequencies ranging from 4 -9 MHz.

PATIENT PREPARATION:

- Introduce yourself to patient

- Verify patient identity according to hospital procedure

- Explain the test

- Obtain patient history including symptoms

- Place the patient in a supine position

GENERAL GUIDELINES:

- A complete examination includes evaluation of the entire course of the accessible portions of each vessel.

- Bilateral testing is considered an integral part of a complete examination.

- Limited examinations for recurring indications may be performed as noted.

- Variations in technique and documentation for assessment of peripheral vascular interventions (i.e., stents), must be described.

TECHNIQUE:

- Equipment gain and display settings will be optimized while imaging vessels with respect to depth, dynamic range and focal zones.

- Color flow Doppler will be added to supplement B-mode images with proper color scale to demonstrate areas of high flow and color aliasing.

- Power Doppler will be used to validate low flow states or occlusions.

- Cursor sample size will be small and positioned parallel to the vessel wall and/or direction of blood flow.

- A spectral Doppler angle of 60 degrees or less will be used to measure velocities.

- Spectral Doppler gains will be set to allow a spectral window and optimized to reduce artifact.

- Areas of suspected stenosis or obstruction will include spectral Doppler waveforms and velocity measurements recorded at and distal to the stenosis or obstruction.

- Sites of intervention (i.e., stents) will include spectral Doppler waveforms and velocity measurements from the proximal, mid and distal sites.

- Plaque should be assessed and characterized.

DOCUMENTATION:

- Duplex evaluation is performed bilaterally starting with the right side.

- Long axis B-mode images must be obtained from:

  - Common carotid artery (CCA)

  - Carotid artery bifurcation (if both the internal carotid artery (ICA) and the external carotid artery (ECA) are not visible in one image, take an image of each with the bifurcation)

  - Internal carotid artery (ICA)

- Spectral Doppler waveforms and velocity measurements must be documented from:

  - Proximal CCA peak systolic velocity (PSV)

  - Mid CCA PSV

  - Proximal ICA PSV and end diastolic velocity (EDV)

  - Mid ICA PSV and EDV if stenosis is present

  - Distal ICA PSV and EDV

  - ECA PSV (temporal tap may be used to identify)

  - Vertebral Artery PSV

- The ICA/CCA ratio should be calculated using a normal CCA PSV and the highest ICA PSV.

- If a stent is present, the following long axis B-mode images must be documented from:

  - Proximal and distal ends

- If a stent is present, spectral Doppler waveforms and velocity measurements must be documented from:

  - native artery at the proximal end of the stent;

  - proximal stent;

  - mid stent;

  - distal stent;

  - native artery at the distal end of the stent.

PROCESSING:

- Review examination data and process for final interpretation

- Note study limitations