Sample Protocol for Performing Lower 
Extremity Arterial Duplex Examinations

Purpose: Duplex examinations are performed to provide evaluation of the lower extremity arteries to assess for plaque morphology location and severity.

INDICATIONS:

Common indications:

- Claudication

- Rest pain

- Follow up of known stenosis

- Ulceration

- Post-op or post-intervention

- Trauma

CONTRAINDICATIONS:

- Patients with casts or bandages

EQUIPMENT:

- Duplex ultrasound with color flow Doppler with transducer frequencies ranging from 3.5-10 MHz.

PATIENT PREPARATION:

- Introduce yourself to patient

- Verify patient identity according to hospital procedure

- Explain the test

- Obtain patient history including symptoms

- Place the patient in a supine position

GENERAL GUIDELINES:

- A complete examination includes evaluation of the entire course of the accessible portions of each vessel.

- Bilateral testing is considered an integral part of a complete examination.

- Limited examinations for recurring indications may be performed as noted.

- Variations in technique and documentation for assessment of peripheral vascular interventions (i.e., stents), must be described.

TECHNIQUE:

- Equipment gain and display settings will be optimized while imaging vessels with respect to depth, dynamic range and focal zones.

- Color-flow Doppler will be added to supplement B-mode images with proper color scale to demonstrate areas of high flow and color aliasing.

- Power Doppler will be used to validate low flow states or occlusions.

- Cursor sample size will be small and positioned parallel to the vessel wall and/or direction of blood flow.

- A spectral Doppler angle of 60 degrees or less will be used to measure velocities.

- Spectral Doppler gains will be set to allow a spectral window and optimized to reduce artifact.

- Areas of suspected stenosis or obstruction will include spectral Doppler waveforms and velocity measurements recorded at and distal to the stenosis or obstruction.

- Sites of intervention (i.e., stents) will include spectral Doppler waveforms and velocity measurements from the proximal, mid and distal sites.

- Plaque should be assessed and characterized.

DOCUMENTATION:

- Duplex evaluation is performed bilaterally starting with the right side.

- Long axis B-mode images must be obtained from:

  - Common Femoral Artery (CFA)

  - Profunda Femoris Artery (PFA)

  - Superficial Femoral Artery (SFA)

  - Popliteal Artery

  - Aorta, common and external iliac arteries and tibial arteries when appropriate

- Spectral Doppler waveforms and velocity measurements must be documented from:

  - CFA

  - PFA

  - SFA

  - Popliteal Artery

  - Posterior Tibial Artery (PTA)

  - Anterior Tibial Artery (ATA)/ Dorsalis Pedis Artery (DPA)

  - Peroneal Artery

  - Aorta, common and external iliac arteries when appropriate

- If a stent is present, the following long axis B-mode images must be documented from:

  - Proximal and distal ends

- If a stent is present, spectral Doppler waveforms and velocity measurements must be documented from:

  - native artery at the proximal end of the stent;

  - proximal stent;

  - mid stent;

  - distal stent;

  - native artery at the distal end of the stent.

- An Ankle/Brachial Index (ABI) must be performed.

  - Bilateral brachial artery systolic pressures must be obtained from both arms and the higher of the two pressures used to calculate the ABI.

  - Ankle systolic pressures must be obtained bilaterally from the distal PTA and ATA/DPA and the higher of the two pressures on each side used to calculate the ABI.

  - Calculate the ABI by dividing the highest ankle systolic pressure from each leg by the highest brachial systolic pressure.

PROCESSING:

- Review examination data and process for final interpretation

- Note study limitations