Sample Protocol for Performing Lower Extremity Reflux Examinations for Venous Insufficiency

Purpose: Peripheral venous reflux examinations are performed to assess the deep and superficial venous system for patency and competency.

INDICATIONS:

Common indications:

- Swelling

- Pain

- Ulcer

- Discoloration

- Visible veins

CONTRAINDICATIONS:

- Patients with casts, bandages

- Patients who are unable to cooperate due to mental status changes or involuntary movements

- Patients with severe edema

EQUIPMENT:

- Duplex ultrasound with color flow Doppler with transducer frequencies ranging from 4 -9 MHz

PATIENT PREPARATION:

- Introduce yourself to patient

- Verify patient identity according to hospital procedure

- Explain the test

- Obtain patient history including symptoms

- Place the patient in a standing, sitting or reverse Trendelenburg position

GENERAL GUIDELINES:

- A complete examination includes evaluation of the entire course of the accessible portions of each vein

- Bilateral testing will be performed unless the patient has unilateral symptoms

- Limited examinations for recurring indications may be performed as noted

- Variations in technique and documentation for assessment of peripheral vascular interventions must be described.

TECHNIQUE:

- Equipment gain and display settings will be optimized while imaging vessels with respect to depth, dynamic range and focal zones

- Spectral Doppler waveform assessment will be done in long axis and will be displayed below the baseline

- Spectral Doppler waveform will be assessed for retrograde flow and measured with calipers on the spectral Doppler waveform.

- Transverse grayscale imaging will be performed with and without transducer compressions

- Diameter measurements will be obtained in transverse plane

- The entire length of the veins will be evaluated

DOCUMENTATION:

- Transverse grayscale images with and without compression must be obtained from:

  - Common femoral vein

  - Saphenofemoral junction

  - Proximal femoral vein

  - Mid femoral vein

  - Distal femoral vein

  - Great saphenous vein

  - Popliteal vein

  - Small saphenous vein

- Spectral Doppler waveforms must be obtained from:

  - Common femoral vein

  - Saphenofemoral junction

  - Great saphenous vein at proximal thigh

  - Great saphenous vein at knee

  - Great saphenous vein below knee

  - Femoral vein at mid thigh

  - Popliteal vein

  - Anterior accessory saphenous vein (when identified)

  - Small saphenous vein at the junction of the deep system (when visualized)

  - Small saphenous vein at mid calf

  - Perforator vein waveforms with active or healed ulcer, as required by the protocol

  - Additional waveforms as required by the protocol

- Transverse grayscale diameter measurements must be obtained from:

  - Saphenofemoral junction

  - Great saphenous vein at proximal thigh

  - Great saphenous vein at knee

  - Anterior accessory saphenous vein (when identified)

  - Small saphenous vein (at saphenopopliteal junction, if visualized)) If not visualized there, the small saphenous vein at the mid calf must be documented.

PROCESSING:

- Review examination data and process for final interpretation

- Note study limitations